## Slide 1
ITk Common Environmental Monitoring System:
Radiation Monitoring
I. Mandić et al.
Jožef Stefan Institute, Ljubljana, Slovenia

## Slide 2
Locations
fixed on the blades of disk 1 or disk 2
inside tracking volume, connection goes
through the enclosure
fixed on outer side of Outer Cylinder
strip EC service volume, outer side of
the bulkhead
Pixel PP1 external volume
position/detector volume | Number of sensors
Strip barrel outer cylinder, z ~ 0, r = 1100 mm, one in each quadrant | 4
Strip EC volume (r ~ 600 mm, z ~ 1600 mm in each quadrant, on side A and C) | 8
Strip EC PP1 volume, fixed on the outer side of structural bulkhead: r ~ 800 mm, z ~ 3000 mm, side A and C, one in each quadrat | 8
Pixel service volume, fixed on outer side of the enclosure r ~ 500 mm, z ~ 3200 mm, side A and C, one in each quadrant | 8
Total | 28

## Slide 3
Radiation Monitor Sensor Board (RMSB)
example from ATLAS ID shown on photos
Very similar design for ITk:
3 RadFETs (TID)
temperature sensor (10 kOhm NTC)
3 BPW34 diodes (NIEL)
Pigtail cable
Sensor side
Back side
( heater for temp stabilization)
PCB
Ceramic (Al2O3)
front side: sensors
back side: resistive layer
~ 2.2 cm

## Slide 4
Sensor box
3 cm x 3 cm
8 mm thick
Photo of the connector on the mockup :
ERNI, 1.27 mm pitch, flat cable, 30 AWG wires
this is an example of a possible cable
and connector
Mockup with ~ 1 m pigtail
hybrid fixed inside PEEK box, 3 cm x 3 cm, 0.8 cm thick,
cable soldered and clamped to the PCB
 can be standard flat cable or e.g. copper clad aluminium wires as used in the ID
Radiation Monitor Sensor Board (RMSB)
RMSB with pigtail, mount to support and connect via connector?
for locations in the tracking volume: solder to prepared cables at assembly site (e.g. NIKHEF)?

## Slide 5
Grounding and shielding requirements (Alex Grillo)
Type 0 (flat cable)
Type II DCS
Type III DCS
USA15
Floating power supply
(DAC)
Cable shield
Return wire
RMSB shield
RMSB outside of the tracking volume:
Connect to faraday cage ground at RMSB mounting location
RMSB inside tracking volume:
Connect to ground at PP1 through shield around Type 0 cable
Ground
PP1
PP2
sensor
Not known:
flat cable shield?
 wrap the cable with foil?
RMSB shield (25 um Al sufficient)?
 wrap PEEK box? metalize PEEK box?
make box from Al not PEEK?

## Slide 6
Readout
readout based on ELMB2  same functionality as ELMB, communication with DCS via ELMB2
radiation monitor readout system running on ELMB in the ID  adapt ID radmon firmware for ELMB2
 build new current source unit controlled by ELMB2:
a) chose DAC compatible with ELMB2 firmware
b) output up to 100 V
 readout system must meet grounding and shielding requirements
 floating power supply referenced at the Faraday cage
 design of current sources and monitoring circuits will follow the solutions described in:
ATL-IP-ES-0106 The Supply and Control system for the Opto Link of the ATLAS pixel detector
Read sensors one by one,
few times per hour
Connected to F.C.
Would current
monitoring still
work?