## Slide 1
Measurements with Si detectors irradiated to extreme fluences
V. Cindro, A. Gorišek, B. Hiti, G. Kramberger, I. Mandić, M. Mikuž, P. Skomina, M. Zavrtanik et al.
Jožef Stefan Institute, Ljubljana, Slovenia
University of Ljubljana, Slovenia

## Slide 2
Introduction
presentation at 34th RD50 Workshop in Lancaster, June 2019
  https://indico.cern.ch/event/719814/contributions/3022499
Charge collection and TCT measurements with 50 µm LGAD irradiated to extreme fluences
 unusual behavior observed in device irradiated to 1e17 n/cm2
 transition from “normal” to “LGAD” kind of behavior (charge multiplication at low bias voltages) after
few weeks at room temperature
this contribution: 75 µm (epitaxial) LGAD irradiated up to 1e17 n/cm2
 CC and TCT measurements
 annealing at 60°C up to 1200 minutes

## Slide 3
Samples
3.3 mm
CNM LGAD run 9256
75 µm epitaxial on low resistivity Cz substrate
chip thickness ~ 600 µm
Irradiated with reactor neutrons to 2.5e16, 5.7e16 and 1e17 n/cm2
Edge-TCT: depletion depth, active thicknes
Q-TCT: charge collection, I-V
Measurements
Plan
systematically measure the effect of annealing time on these diodes
annealing times: 0, 80, 240, 560 and 1200 minutes
measure in reverse and forward bias

## Slide 4
Guard ring
Pad electrode
10 V
Guard ring
Pad electrode
20 V
No signal on pad electrode
before depletion of multiplication layer
1/C2
Bias (V)
Pin diode
LGAD
M. Carulla, TREDI 2017, Trento
Edge - TCT
Signal from guard ring or pad electrode
IR laser direction
Not irradiated
Not irradiated
Diode Surface
depth

## Slide 5
Ф = 0
Bias = 200 V
Pad electrode
Guard ring
Field at the back,
effect enhanced because of large pad
electrode.
 Measure depletion depth on guard ring
Depletion
depth
Edge- TCT

## Slide 6
Neff = 1.8e14 cm-3  Resistivity ~ 70 Ωcm, Expected Vfd (75 µm) ~ 750 V
Nominal: 100 Ωcm  Neff = 1.3e14 cm-3  Vfd ~ 550 V
Edge TCT, before irradiation
Signal from guard ring

## Slide 7
Forward bias
Forward bias
Forward bias
charge from the 75 um epitaxial layer, no charge from low resistivity substrate
electric field across whole detector thickness already at 100 V in reverse or forward bias
E-TCT, irradiated detectors

## Slide 8
E-TCT
Detector surface
Depth
Uniform response across x coordinate
Ф = 1e17 n/cm2, Bias = 600 V, signal from guardring
IR laser direction
x

## Slide 9
detector
ORTEC 142 charge sensitive amplifier + custom made shaper with 25 ns shaping time
digital oscilloscope
Charge Collection measurement with Sr-90
Large (3x3 mm) pad detector and small collimator holes enable recording of a clean sample of waveforms
 no “empty triggers”  can measure average charge collected after passage of a (almost) MIP
also at small signal/noise
New since last year: a) replaced bias resistor in ORTEC 142  1 MΩ instead of 10 MΩ
b) improved cooling (measure at -30 C)
Can measure up to higher
bias voltages

## Slide 10
75 μm epitaxial LGAD
M. Carulla, TREDI 2017, Trento
Measured with Sr90,
CCE = measured_mean/7500 el
LGAD 75 um, relatively low gain
Measured with alpha source:

## Slide 11
G. Kramberger et al, 2013 JINST 8 P08004
“Magic formula” (300 µm spaghetti Ф > 1e15):
QMPV = k∙Фb∙V,
k = 26.4 el/V, b = -0.683
(Ф in 1e15 n/cm2, V in volts)
More charge at same bias voltage for thinner detectors
magic formula slightly modified for
75 µm detectors:
Qmean_75um = k∙Фb∙V
k = 44 el/V, b = -0.56
(Ф in 1e15 n/cm2, V in volts)
Charge collection
Annealed 80 minutes
Dots: measurements

## Slide 12
Compare with 50 µm LGAD
50 µm epitaxial LGAD (see https://indico.cern.ch/event/719814/contributions/3022499 )
no significant differences:
  50 µm at 2.8e16 slightly more charge than 75 µm at 2.5 e16
  not much difference at 1e17
   can be approximated with the same magic formula
Dots: measurements 75 µm

## Slide 13
Charge collection – compare with noise
not useful to bias far above 700 V
  start of breakdown: current increase, breakdown noise

## Slide 14
Forward bias
For lower fluence at lower forward bias voltages somewhat more charge than in reverse bias
at 1e17 not much difference
lower current in reverse bias  can go to higher bias voltages
 better performance under reverse bias
 difference smaller at higher fluence
Dots – reverse bias
Squares – forward bias
Dots – reverse bias
Squares – forward bias

## Slide 15
forward bias
reverse bias
All measurements, different annealing times
Charge collection - annealing
Not irradiated
2.5e16
5.7e16
1e17

## Slide 16
Annealing: collected charge
Forward bias
 not much annealing

## Slide 17
2.5e16
5.7e16
1e17
Forward bias
Reverse bias
T = - 30° C
reverse bias:
current increases with increasing fluence
 increase of generation current
forward bias:
current decreases with increasing fluence
 increase of resistance because of
less free carriers, decrease of mobility
Current
Diode like IV also after 1e17, breakdown above 700 V
Annealed 1200 min

## Slide 18
Ф = 2.5e16, I = αФV = 12 µA
Ф = 5.7e16, I = αФV = 28 µA
Ф = 1e17, I = αФV = 49 µA
T = - 30° C
E-TCT: field in the whole detector volume at V > 100 V  is current generated in the whole detector volume?
current increases with bias
above ~600 V current higher than calculated from α = 4e-17 Acm-1  charge multiplication
 noise in CC measurement starts to increase sharply above ~ 600 V  micro-discharges
 contribution of surface or edge currents?
Current

## Slide 19
Annealing of current (reverse bias)
roughly follows the shape of annealing curve:
RD48 values: α1 = 1e-17 Acm-1, t1 = 93 min, α0 = 5e-17 Acm-1, ẞ =3.3e-18 Acm-1
 annealing during irradiation
(irradiation time 4 h, T > 40° C,
temp. probably higher for higher fluences)
Forward bias
 annealing effects small
Points at t = 80 min set by hand to match data

## Slide 20
Summary
measurements with 75 um thick epitaxial LGAD irradiated 2.5e16, 5.7 e16 and 1e17
E-TCT:
  before irradiation: as expected, observed LGAD depletion behavior observed
  after irradiation response across whole detector thickness already at 100 V in reverse or forward bias
  no signal from low resistivity substrate
Charge collection measurements with Sr-90:
  charge scales with fluence and voltage according to Qmean_75um = k∙Фb∙V (k = 44 el/V, b = -0.56, Ф in 1e15 n/cm2, V in volts)
  at high bias voltage sharp increase of noise because of micro-discharges
   detector can be operated up to ~ 600 V reverse bias, ~ 2000 el (mean) collected at 1e17
  better charge collection in reverse bias than in forward bias
   differences get smaller at higher fluences
  change of collected charge by annealing at 60°C up to 1200 minutes not very significant in these measurements
Current
  reverse current increases with fluence, forward current drops
  reverse bias: linear increase with bias voltage up to ~ 600 V, sharp increase after that
  higher forward than reverse current, differences get smaller with increasing fluence