## Slide 1: CCE measurements on ATL17LS ministrip sensors in Ljubljana –comparison with R0, thin and thick active regions - preliminary results!!!
Sensors irradiated at TRIGA reactor in Ljubljana, 1 Mev equivalent fluences 1014, 5 1014, 1015, 2 1015 cm-2
measurements at T = -20C
standard cuts – seed 3.5 σ, neighbour 1.5 σ
“narrow” collimation
“narrow “ clusters
thin and thick sensors (nominal active thickness 240 μm and 290 μm) with same fluence measured on the same Alibava board
Cluster width 300 V, neutrons 5 1014 cm-2

## Slide 2: Fluence 1014 ncm-2
R0 fluence = 0

## Slide 3
Fluence 5 1014 cm-2
Voltage on ATL17LS limited by breakdown on supporting board

## Slide 4: Fluence 1015ncm-2
RO has higher leakage current – higher fluence ?

## Slide 5: Fluence 2 1015 ncm-2

## Slide 6: Conclusions
ATL17LS and R0 have same CCE within the measurement uncertainty
discrepancy at 1015cm-2not clearly understood, different fluences?
thin and thick sensors have very similar CCE after irradiation to 5 1014 cm-2 and more
thin sensor has smaller signal at the low fluence compared to standard – as expected for 240 μm thick active region