## Slide 1
TPA-TCT measurements status
F9 Weekly, 06. 08. 2021
Bojan Hiti, F9, Jožef Stefan Institute (JSI)

## Slide 2: Ljubljana TPA-TCT setup
Based on FYLA LFC 1500X fs laser (λ=1550 nm)
Readout electronics and DAQ based on standard Particulars TCT

## Slide 3: Sample and measurement
Sample:
  6 x 6 mm p-in-n pad diode, metalization opening 2 mm
  Thickness 300 μm, Vfd=20 V, 10 kΩ·cm, unirradiated
Charge = time integral (25 ns) of TCT-pulse
Amplifier output calibrated with Am241 alpha source (5.4 MeV)
Beam
z
x
Direction of view

## Slide 4: Sample tilt
Sample is tilted by approx. 1°
Ordered a manual tilt stage to compensate

## Slide 5: Beam characterization

## Slide 6: Beam characterization
Fit Error function at different z
extract w(z)
extract w0

## Slide 7: Beam characterization

## Slide 8: Beam characterization
z0 = 18 um
signal size

## Slide 9: Longitudinal spot size
Errata z->x

## Slide 10: Longitudinal spot size
Reducing beam intensity (approx. factor 2 in charge)
Cigar half-length cca 15 um

## Slide 11: SEU potential
LGAD mortality studies require charge of the order of 30M e-h pairs
  Can it be done with TPA?
Maximized laser signal
HPK LGAD 2.5e15 biased to 700 V  failed in DESY test beam
No defects observed with TPA
There is still potential to increase pulse energy, but setup overhaul
  Reduce losses on optical bench
  Pin = 100 mW, Pout = 15 mW

## Slide 12: Edge TPA-TCT
 | Strip size 630 x 40 um | 
Try Edge-TCT configuration using strip detector (CHESS 2)
So far no signals yet observed

## Slide 13: BACKUP

## Slide 14
FYLA Laser Unit
Pulse management module
(trigger & pulse energy)
Pulse compressor
Slide-in autocorrelator
XYZ sample placement
Optional cooling
Sample housing
Objective NA=0.65,
Working distance 10 mm
Particulars
readout system

## Slide 15: TPA laser pulse: Time domain
Measurement of laser pulse duration with an autocorrelator
Autocorrelator output
Pulse width (FWHM)
0.196 ms x 1600 fs/ms = 310 fs
Pulse duration ≈ 300 fs