## Slide 1
First CERN SPS test beam with BCM' and Calypso 3
F9 Weekly, 27. 08. 2021
Bojan Hiti, F9, Jožef Stefan Institute (JSI)

## Slide 2: Samples
Calypso 3 chip bonded to segmented diamond samples (2 boards):
  Board 1: 3d diamond with 4 test pads (50 V)
    400x400 um with bias grid P0
    200x200 um with bias grid P1
    400x400 um without bias grid P2
    200x200 um without bias grid P3
  Board 2: Planar single crystal single pad diamond (200 V)
Each channel provides:
  Analog output after first stage amplifier
  (nominal < 1 ns rise time, ~5 ns fall time)
  Two digital comparator outputs (LVDS+, LVDS-)
Read out with DRS4, 3 inputs at a time
Signal  Integration window 20 ns around signal peak (from avg. waveform)

## Slide 3: Beam and telescope
SPS 120 Gev pion beam
MALTA CMOS telescope
  6 planes (3__0.5m_DUT_0.5m__3)
  36x36 um pixels, configurable rectangular Region of Interest
  Synchronous trigger signal – jitter 10 ns  not optimal
  Scintillator upgrade is ongoing

## Slide 4: Run log summary
Run number | Run summary | N events | Comment
634 | P0 Analog out | 50k | 
637 | P1 Analog out | 50k | 
641 | P2 Analog out | 2M | device not working/connected
651 | P0 comparator thr. 0x70 (high) | 20k | 
652 | P0 comparator thr. 0x30 (medium) | 20k | 
653 | P0 comparator thr. 0x10 (low – operation?) | 40k | 
670 | P0 comparator thr. 0x10 (low), -50 V | 80k | 
674 | Planar SC 2x2 mm | 500k | 

## Slide 5: Analog out P0 and P1 (run 637)

## Slide 6
Threshold scan (high, medium, low)

## Slide 7: Run 651 (P0 high threshold)
Analog
Discriminator

## Slide 8: Run 652 (P0 medium threshold)
Analog
Discriminator

## Slide 9: Run 653 (P0 low threshold)
Analog
Discriminator

## Slide 10
Negative Vbias = -50 V
taken immediately after changing voltage – no polarization effects

## Slide 11: Run 670 (low threshold)
Analog
Discriminator

## Slide 12: Run 670 cross check

## Slide 13
Single crystal pad diamond
Medium threshold
ROI smaller than sample ...

## Slide 14: Medium threshold
Analog
Discriminator