## Slide 1: Develop Lab conditions (radiation protection) needed for producing the isotopes
Athanasios Zacharakis
Medical Physicist/Radiation Protection Officer
BIOKOSMOS SA

## Slide 2
PURPOSE AND RESEARCH SCOPE
Designing of a radio-isotope laboratory
Safety requirements for chemical processing of 176Yb irradiated targets.
Focuses on Hot Laboratory, QC Laboratory, and Waste Management Area.
Compliance with Greek Atomic Energy Commission, Greek National Organization of Medicines and International Atomic Energy Agency basic safety standards.
Aims for occupational doses below 6 mSv/year design limit.

## Slide 3
SHIELDING DESIGN: PRIMARY BARRIERS
Shielding Zone | Critical Target Energy | Structural Specification
Target Receipt & Decay Vault | 1.24 MeV Gamma (177Yb) | 1.6 m High-Density Concrete Barrier
Processing Hot Cells | 208 keV Gamma (177Lu) | 75 - 100 mm Lead (Pb) Enclosure
QC Laboratory Benches | Low Energy Aliquots | 20 mm Pb Shielded L-Blocks
Radiochemistry Fume Hoods | Bremsstrahlung / Gammas | 50 mm Pb for the base and side walls with 75 mm vertical sliding Pb Glass
Mandatory 24-hour cooling period for the targets. The presence of 177Yb introduces a transient but severe high-energy gamma hazard in the first hours after EOB.

## Slide 4
RADIATION MONITORING PROGRAM
Comprehensive radiation monitoring framework is essential for validation of shielding structure performance and early detection of containment failures.
Fixed area radiation monitors installed in Hot Lab, QC, and Waste Room
Continuous exhaust stack tracking for particulate and volatile releases
Portable radiation detectors
Personal dosimetry includes whole-body badges, ring dosimeters (exchange monthly) and electronic personal dosimeters (provide real-time dose alarms)

## Slide 5
WASTE MANAGEMENT: Strategy and Steams
Stream A: Target Raffinate
Contains stable 176Yb and radioactive 175Yb isotopes.
Stored for minimum 42 days for decay and recycling.
Stream B: Operational Waste
General contaminated consumables, gloves, and empty vials.
Stored in drums for minimum 67 days.
Released for conventional disposal only after clearance (< 0.1 μSv/h).
The waste management infrastructure will be designed to allow safe handling, storage, and conditioning of all radioactive materials produced by irradiation of Ytterbium targets. The strategy includes radiological protection, chemical safety, and environmental considerations focusing on: Management of short-lived beta-emitting radionuclides (177Lu and residual Yb isotopes), handling of chemically aggressive liquid waste streams (acidic media), minimization of waste volumes and promotion of target material recovery and prevention of airborne contamination and radiolytic gas accumulation.

## Slide 6
Safety Protocols & Emergency Response
Contingency Planning:
  “Stop and shield, Warn, Isolate, Monitor, Survey and decontamination” (SWIMS) protocol for an unsealed radiological release.
  Strategic placement of emergency spill kits.
  Designated decontamination zones with dedicated plumbing containment.
Immediate Actions:
  Evacuation of work area during high-level alarms.
  Localized isolation of automated hot cells via panic buttons.
  Reporting chain-of-command for radiological contamination incidents.

## Slide 7
STAFF TRAINING PROGRAM
At the end of the program, staff should be able to:
Explain the basic physics, biological effects, and specific risks of 177Lu and associated isotopes (177Yb, 175Yb)
Apply ALARA principles (time, distance, shielding) in routine tasks
Operate hot cells, L-blocks, and remote tools without unnecessary exposure
Identify and control contamination risks in hot lab, QC, and waste areas
Interpret area monitor and survey meter readings and act appropriately
Respond correctly to spills, abnormal dose-rate readings, and alarms
Document activities, surveys, and incidents in compliance with regulations

## Slide 8
Conclusions
Technical Summary
Shielding Efficacy:
  Mandatory 24-hour cooling period for the targets
  Shielding must account for 1.2 MeV transient 177Yb emissions.
  Lead thickness of 75-100 mm ensures safety during 177Lu processing.
  Safety Integration:
  Real-time area monitoring provides reliable detection of radiological anomalies.
Operational Summary
Strategic Compliance:
  Design complies with ALARA and 6 mSv/year design limits.
  Clearance protocols ensure safe, documented environmental waste disposal.

## Slide 9: Thank you for your attention